Typical Temperature Calibration for GSL1600X-80 Tube Furnace

In order to achieve accurate firing temperature, you shall calibrate temperature between inside tube and outside tube before using by yourselves via NIST certified precision temperature calibrator  ( click to order one) and thermocouple.

Please be advised that GSL1600X-80 alumina tube furnace is designed for use above 1400oC.  Under 1400oC,  temperature inside is higher than outside ( controller’s).  The below table show a typical calibration curve of temperature difference.

- You shall set temperature off-set  ( e.g. set SC value in controller, please read manual to learn how )

- The furnace is shipped without temperature off-set.

- The tube furnace ;s constant temperature zone length  (+/-1oC) at  > 1400oC is about 150 mm in center area.  You shall place your sample into the zone.

- MTI is not responsible for any damage caused by improperly using the furnace

If you have any question for operation the furnace, please feel free to contact us by email at info@mtixtl.com

# MTI Corporation

860 South 19th street, Richmond, CA 94804, USA

tel: 510-525-3070, fax: 510-525-4705 email: info@mtixtl.com

www.mtixtl.com